## Slide 1
Saintgits College of Engineering (Autonomous)
NON-INVASIVE ANEMIA DETECTOR
Project In Charge: Er. Anish Thomas (Asst Prof.)
Project Guide : Er. Jyothish Chandran G (Asst Prof.)
Presented By:
Britty Siby(MGP23EC024)
Devika S (MGP23EC026)
Devika S Nair (MGP23EC027)
Swetha Anil(MGP23EC066)

## Slide 2: CONTENTS
Saintgits College of Engineering (Autonomous)
Introduction
Initial Scope of Work
Changes in plan
Literature Survey
Hardware Components
Block Diagram
Working Principle
Algorithm
Calibration graph
Results
Functional prototype
Future Scope
Acknowledgement
References
Conclusion

## Slide 3: INTRODUCTION
Saintgits College of Engineering (Autonomous)
Anemia is a widespread health condition.
Conventional diagnosis requires blood samples and uses invasive methods.
This brings us to the idea of non-invasive, quick, and cost-effective solution.
Proposed system uses optical sensing (Red & IR light)
Using lights of different wavelengths to estimate haemoglobin levels.

## Slide 4: INITIAL SCOPE OF WORKS
Saintgits College of Engineering (Autonomous)
Measure haemoglobin non-invasively
Use MAX30102 sensor for data collection
Develop a simple prototype using ESP32
Display results on OLED
Provide quick estimation for demonstration

## Slide 5: CHANGES IN PLAN
Saintgits College of Engineering (Autonomous)
1. Improved to multiple readings + averaging
2. Modified to timed sampling (~7 seconds) for stability
3. Updated to button-controlled measurement system
Single reading approach
Instant measurement
Continuous sensing

## Slide 6: LITERATURE SURVEY
Saintgits College of Engineering (Autonomous)
S.no. | Reference Paper | Information gained
1 | Non-invasive portable hemoglobin concentration monitoring system using optical sensor for anemia disease | This paper presents a portable, non-invasive system to measure hemoglobin using red and infrared light. The device analyzes light absorption through tissue and uses calibration with lab data to estimate Hb levels. It enables quick and painless anemia screening.
2 | Non-invasive optical method for Anemia Detection | This study explores an optical approach for detecting anemia without blood extraction. It relies on analyzing variations in light absorption (red and infrared) to estimate hemoglobin levels.
3 | Estimate of Anemia with new Non-invasive systems a moment of reflection | This paper reviews recent advancements in non-invasive techniques for anemia estimation. It discusses various optical methods and their effectiveness in measuring hemoglobin without blood tests.

## Slide 7: HARDWARE COMPONENTS
Saintgits College of Engineering (Autonomous)
ESP32 Microcontroller
MAX30102 Sensor
OLED Display (SSD1306)
Push Button
USB Power Supply

## Slide 8: BLOCK DIAGRAM
Saintgits College of Engineering (Autonomous)

## Slide 9: WORKING PRICIPLE
Saintgits College of Engineering (Autonomous)
Light is transmitted through finger
Hemoglobin absorbs light differently
Sensor measures:
  Red light
  Infrared light
Ratio is calculated
Ratio correlates with hemoglobin level

## Slide 10: ALGORITHM
Saintgits College of Engineering (Autonomous)
Collect Red & IR sensor values for 7 secs
Average multiple readings
Compute ratio:
R = Red / IR
Use calibration formula for Hb estimation:
Hb= -148.25 R + 120.98
Display result on OLED

## Slide 11: CALIBRATION GRAPH

## Slide 12: RESULTS
Saintgits College of Engineering (Autonomous)
The prototype successfully acquired the data from the finger placed on the sensor for the specified time 7 seconds for calculation of Hb level
Calculated the Hb levels based on the data i.e. amount of reflected red and infrared light captured by the MAX30102 sensor.
Showed the approximate Hb level of the individuals
Classified the result as : Normal if Hb>12g/dL
Mild anaemia if 10g/dL<Hb<12g/dL
Anaemic if Hb< 10g/dL

## Slide 13: THE FUNCTIONAL PROTOTYPE
Saintgits College of Engineering (Autonomous)

## Slide 14: FUTURE SCOPES
Saintgits College of Engineering (Autonomous)
Improve accuracy through clinical calibration with larger datasets from laboratory
Incorporate machine learning models for better Hb prediction
Introduce gender-based calibration (male/female separation) for improved precision
Enhance measurement using skin texture and skin tone analysis
Integrate color detection techniques to reduce signal variation
Develop a mobile application for real-time monitoring

## Slide 15: CONCLUSION
Saintgits College of Engineering (Autonomous)

## Slide 16: ACKNOWLEDGMENTS
Saintgits College of Engineering (Autonomous)
Dr. Shiny G -HOD-ECE & AE
Er. Jyothish Chandran G- Assistant professor(Guide)
Er. Anish Thomas-Assistant professor(Project Coordinator)
Er. Melvin C Jose-Assistant professor(ECE)

## Slide 17: REFERENCES
Saintgits College of Engineering (Autonomous)
Chuchart Pintavirooj, Baorong Ni, et al., “Non-invasive portable hemoglobin concentration monitoring system using optical sensor for anemia disease,” 2021.
P. Chamola, Poornima Mittal, “Non-invasive optical method for Anemia Detection,” 2023.
Giovanni Dimauro, Serena De Ruvo, et al., “Estimate of Anemia with new Non-invasive systems: a moment of reflection.”
MAX30102 Sensor Datasheet, Maxim Integrated.
Research articles on Near Infrared (NIR) Spectroscopy in biomedical applications.

## Slide 18: THANK YOU
Saintgits College of Engineering (Autonomous)
1